## Slide 1: Accessible Medical Diagnostic Equipment & Prescription Drug Container Labels
Pacific ADA Center webinar
3/26/2020

## Slide 2: Agenda
Medical Diagnostic Equipment
  Background
  Ch 1: Application
  Ch 2: Scoping
  Ch 3: Technical requirements
Prescription drug containers

## Slide 3: Medical Diagnostic Equipment

## Slide 4: Background

## Slide 5: Existing Standards
ADA and ABA Accessibility Guidelines
ANSI/AAMI HE75 – 2009 Human Factors Engineering – Design of Medical Devices
RERC on Accessible Medical Instrumentation
Research Data
Notes: (RERC) Rehabilitation Engineering Research Center on Accessible Medical Instrumentation

## Slide 6: Accessibility challenge
Accessible Medical Instrumentation RERC usability survey
Survey reported usability of various medical diagnostic equipment
Significant percent of respondents reported equipment was moderately difficult to impossible to use
Reports of:
Moderately difficult/impossible to use
Examination tables – 75%
X-ray equipment – 68%
Weight scales – 53%
Examination chairs – 50%
Notes: Rehabilitation Engineering Research Center on Accessible Medical Instrumentation

## Slide 7: Rehabilitation Act §510
Mandates the Access Board to issue minimum technical criteria for accessible medical diagnostic equipment in consultation with FDA
Affordable Care Act amends the Rehabilitation Act to address access to MDE
(March 23, 2010)

## Slide 8: Rehabilitation Act §510
Specifically addresses equipment in:
  Physicians offices, clinics, emergency rooms, hospitals, other medical settings. Includes optometrists and dentists
Scoping to be determined in appropriate regulatory or administrative context
Final Standards are voluntary until adopted by an enforcing Agency e.g., DOJ for the Americans with Disabilities Act
Could be used by FDA to verify labelling claims
Can be voluntarily adopted through policy

## Slide 9: Federal Partners
Health and Human Services
  US Food and Drug Administration
Department of Justice
  Disability Rights Section
Department of Veterans Affairs

## Slide 10: Standards for Accessible Medical Diagnostic Equipment
Published in:
  Code of Federal Regulations, Title 36, §1195.1 Standards, Appendix to Part 1195
Found on our website
  www.access-board.gov
Chapter 1: Application and Administration
  M101 General
  M102 Definitions
Chapter 2: Scoping
  M201 General
Chapter 3: Technical Requirements
  M301 Diagnostic Equipment Used by Patients in Supine, Prone, or Side-Lying Position
  M302 Diagnostic Equipment Used by Patients in Seated Position
  M303 Diagnostic Equipment Used by Patients Seated in a Wheelchair
  M304 Diagnostic Equipment Used by Patients in Standing Position
  M305 Supports
  M306 Communication
  M307 Operable Parts

## Slide 11: MDE Standards
Do not address the accessibility of existing diagnostic equipment
Do not require a percentage of new or replacement equipment to be accessible

## Slide 12: Chapter 1: Application and Administration

## Slide 13: Application
Purpose:
To ensure accessibility to, and usability of the diagnostic equipment by patients with disabilities.
Provide for independent access and use of diagnostic equipment to the maximum extent possible
Applies to new diagnostic equipment, based on the patient positions that the equipment supports, during patient transfer and diagnostic use
Communication features or operable parts – where provided for patient use

## Slide 14: “Diagnostic equipment” Includes:
Examination tables
Examination chairs (including those used for dental or optical exams)
Weight scales
Radiological equipment
  X-ray machines
  MRI, CT Scan
  Mammography equipment
Other MDE

## Slide 15: “Diagnostic equipment” Does not include:
Personal devices
  Blood glucose monitors
Positioning aids
  Wedges
Surgical and medical instruments
  Stethoscopes
  Forceps, scalpels

## Slide 16: Equivalent facilitation
MDE language is different than 2010 ADA
The use of alternative designs or technologies that result in substantially equivalent or greater accessibility and usability than specified in the MDE Standards is permitted.

## Slide 17: Measurements and Dimensions
Based on Adult Dimensions and Anthropometrics
Use of MDE by children is not addressed in these standards
Dimensions not stated as maximum or minimum are absolute
Subject to conventional industry tolerances
Measured in US or metric units, independently

## Slide 18: Definitions
Undefined terms:
Not defined in M102.1 or in regulations or policies issued by an enforcing authority:
Ordinarily accepted meaning in the sense that the context implies.
Defined in section M102.1:
End Transfer Surface
Enforcing Authority
Examination Chair
Imaging Bed
Imaging Equipment with Bores
Medical Diagnostic Equipment (Diagnostic Equipment)
Operable Parts
Side Transfer Surface
Transfer Surface
Defined in regulations or policies issued by enforcing authority

## Slide 19: Chapter 2: Scoping

## Slide 20: M201 General
Scoping to be determined by enforcing authority
  Potential enforcing authorities could include DOJ, HHS, or adoption through policy (like select VAs)
General exception:
  “rare circumstances where compliance would alter diagnostically required structural or operational characteristics of the equipment and would prevent the use of the equipment for its intended diagnostic purpose”
  Diagnostic equipment shall comply to the maximum extent practicable

## Slide 21: Chapter 3: Technical Requirements

## Slide 22: Chapter 3
Supine, prone, or side-lying positions (M301)
Seated position (M302)
Seated in a wheelchair (M303)
Standing position (M304)
Supports (M305)
Communication (M306)
Operable parts (M307)

## Slide 23: Chapter 3
Supine, prone, or side-lying positions (M301)
Seated position (M302)
Seated in a wheelchair (M303)
Standing position (M304)
Supports (M305)
Communication (M306)
Operable parts (M307)

## Slide 24: Supine, Prone, or Side-Lying Position (M301)
Includes:
MRI, CT scan
Examination tables
Hospital stretcher
Does not include:
Examination chairs that recline (covered in M302)

## Slide 25: Supine, Prone, or Side-Lying Position (M301)
Transfer surface
  Height adjustable
  Size
    End transfer surface
    Side transfer surface
  Unobstructed transfer
Supports
  Transfer supports
  Leg supports, where provided
  Head and back support
Lift compatibility
  Clearance in base
  Clearance around base
  Exception: overhead lifts

## Slide 26: Transfer surface
Height adjustable
  Measured from floor to top of uncompressed transfer surface
  High transfer position at 25 inches minimum
  Low transfer position of (17-19 inches) maximum
  At least 4 additional transfer positions located between the low and high, separated by 1 inch minimum

## Slide 27: Transfer surface
Low position sunset clause
Low transfer position height (17” – 19”) expires in 5 years (January 9, 2022)
Research to be done on transfer heights

## Slide 28: Transfer surface
End Transfer Surface
Located at one end of an examination surface that allows transfer at the end and one adjoining side
Side Transfer Surface
Located within the length of the examination surface that allows transfer on two opposing sides

## Slide 29: Transfer surface
Size
Comply with either:
  End Transfer
    17” x 28” min
  Side Transfer
    28” x 28” min
Measured from center points of their opposing sides

## Slide 30: Transfer surface
Size: Exception
Imaging equipment with bores
  Width: 21” min
  Not less than the full width of the examination surface provided

## Slide 31: Transfer surface
Unobstructed transfer
Each transfer surface shall provide two unobstructed sides for patient transfer
  Side transfer = L side + R side
  End transfer = end + (L/R) side

## Slide 32: Transfer surface
Permitted obstructions
Obstructions less than 3” deep
  Extend beyond transfer sides of transfer surfaces
  Do not protrude above the tops of transfer surfaces
Temporary obstructions provided they can be repositioned during transfer

## Slide 33: Supports
Transfer support
  Required at transfer surfaces
  (covered in M305.2)
Leg supports
  Where stirrups are provided, leg supports shall also be provided
  (covered in M305.4)
Head and back support
  Where diagnostic equipment is used in a reclined position
  (covered in M305.5)

## Slide 34: Transfer supports
Location
End Transfer:
  Short side of transfer surface, opposite transfer side
Side Transfer:
  Opposite sides of transfer surface, removable supports
Gap
  1 ½” maximum gap
  Measured horizontally from nearest edge

## Slide 35: Transfer supports
Length
End Transfer:
  Short side of transfer surface, opposite transfer side
  15” minimum
  At least 13 ½” along depth of the transfer surface
Side Transfer:
  Opposite sides of transfer surface, removable supports
  28” minimum

## Slide 36: Transfer supports
Height
During use, tops of Gripping surfaces must be between 6” minimum and 19” maximum
Supports do not need to be horizontal
Horizontal supports not required – gripping surface must be within min and max height

## Slide 37: Transfer supports
Location: Exception
Where the support folds, collapses, or articulates, the transfer support shall be permitted to be located 3” maximum from the edge of the transfer surface

## Slide 38: Transfer supports
Length: Exception
Where side transfer surfaces are part of an articulating surface, the support is permitted to be 15“ long min.

## Slide 39: Transfer supports
Exception: Wide Imaging Exam Surfaces
Where width of imaging bed is greater than 24”
  Transfer supports can be 12” long min.
  During use, tops of gripping surfaces can be 3” min. and 6” max. higher than the top of the transfer surface.

## Slide 40: Portable Lift Compatibility

## Slide 41: Lift compatibility
Clearance in Base
6 inches high minimum
36 inches deep minimum
  Measured from edge of examination surface
  Or extend full depth if less than 36 inches deep
39 inches wide minimum
Equipment permitted within 8 inches maximum of the centerline

## Slide 42: Lift compatibility
Clearance around Base
6 inches high minimum
36 inches deep minimum
  Measured from edge of examination surface
26 inches wide minimum
  Permitted to increase at a rate of 1 inch in width for each 3 inches in depth

## Slide 43: Lift compatibility
Exception: Overhead lifts
Where fixed overhead patient lifts are provided and:
  Use is permitted by enforcing authority
  Equipment clearly labeled as not compatible with portable floor lifts

## Slide 44: Chapter 3
Supine, prone, or side-lying positions (M301)
Seated position (M302)
Seated in a wheelchair (M303)
Standing position (M304)
Supports (M305)
Communication (M306)
Operable parts (M307)

## Slide 45: Seated Position (M302)
Transfer surface
  Height adjustable
  Size
    2 adjoining sides
  Unobstructed transfer
Supports
  Transfer supports
  Leg supports, where provided
  Head and back support
Lift compatibility
  Exception: overhead lifts
  Clearance in base
  Clearance around the base

## Slide 46: Transfer surface
Height adjustable (same)
  Measured from floor to top of uncompressed transfer surface
  High transfer position at 25 inches minimum
  Low transfer position of (17-19 inches) maximum
  At least 4 additional transfer positions located between the low and high, separated by 1 inch minimum

## Slide 47: Transfer Surface
Two Adjoining Sides
Exception: Chairs with Fixed Footrest

## Slide 48: Transfer Surface
Minimum size
How to measure

## Slide 49: Supports
(same as M301)
Transfer supports
  Required at transfer surfaces
  (covered in M305.2)
Leg supports
  Where stirrups are provided, leg supports shall also be provided
  (covered in M305.4)
Head and back support
  Where diagnostic equipment is used in a reclined position
  (covered in M305.5)

## Slide 50: Lift compatibility
(same as M301)
Clearance under base
Clearance around the base
Overhead lift exception

## Slide 51: Chapter 3
Supine, prone, or side-lying positions (M301)
Seated position (M302)
Seated in a wheelchair (M303)
Standing position (M304)
Supports (M305)
Communication (M306)
Operable parts (M307)

## Slide 52: Seated in a Wheelchair (M303)
Wheelchair spaces
  Orientation
  Width
  Depth
    Front or rear entry
    Pass-through entry
    Side entry
    Equipment clearance
      Breast platforms
      Other equipment
  Surfaces
  Edge protection
Entry
  Beveled
  Ramped
Components
  Breast platform adjustability

## Slide 53: Wheelchair spaces
Orientation
Designed so that a patient seated in a wheelchair orients in the same direction that a patient not seated in a wheelchair orients when the MDE is in use

## Slide 54: Wheelchair spaces
Width :
36” wide minimum
Depth:
Front or rear entry: 36” min
Side entry: 60” min
Pass-through: 40” min
  (weight scales)

## Slide 55: Wheelchair spaces
Width and depth based on ADA maneuvering clearances where spaces are confined on 3 sides (ADA - 305.7)

## Slide 56: Wheelchair spaces
The MDE Standard does NOT address clear floor space for transfers

## Slide 57: Wheelchair Weight Scales
Width exception for Raised Platform (M303.2.2x)
Pass-through entry wheelchair space (M303.2.3.2)
May require edge protection (M303.2.6)
Change in level at entry, may require bevels or ramps (M303.3)

## Slide 58: Wheelchair Spaces on Weight Scales - M303
Pass through entry depth and width exception address

## Slide 59: Ramps Entries
Beveled change in level
Running slope Exception:
consistent with ADA Standards exception for rises of 3” or less in alterations, except max height is 2 ½”
Combination of height and
steep slope a barrier

## Slide 60: Clearances beneath Components (M303)
Most MDE knee and toe clearance will be the same as ADA

## Slide 61: Breast platforms
Height adjustable platform (M303.4.1)
continuously adjustable
low height of 26 inches
high height of 42 inches

## Slide 62: Breast platforms
Clearances beneath Components (M303)
Knee clearance 17” deep min to allow user to bring chest in contact with platform
Knee and toe clearance allows additional obstruction under foot plates

## Slide 63: Chapter 3
Supine, prone, or side-lying positions (M301)
Seated position (M302)
Seated in a wheelchair (M303)
Standing position (M304)
Supports (M305)
Communication (M306)
Operable parts (M307)

## Slide 64: Standing Position (M304)
  Slip resistant standing surface
  Supports provided on two sides
  Exception: Equipment with pass-through entry/exit can provide one horizontal support

## Slide 65: Standing Supports – M305.3
Horizontal Supports:
On most equipment 4” min long
34” min. – 38” max.
Above standing surface to top
Vertical Supports:
18” min. long
34” min. - 37” max.
Above standing surface to lowest end of the gripping surface

## Slide 66: Standing Supports (M305.3)
Horizontal Supports on weight scales with wheelchair spaces:
One entry – gripping surface along 80% of platform length
Pass- through:
  Gripping surface equal to length of platform length
  On one side only (M304.2.2 exception)

## Slide 67: Chapter 3
Supine, prone, or side-lying positions (M301)
Seated position (M302)
Seated in a wheelchair (M303)
Standing position (M304)
Supports (M305)
Communication (M306)
Operable parts (M307)
PUSH

## Slide 68: Transfer and Standing Supports
Structural strength
Not addressed like in ADA and ABA standards for grab bars
More variables
Industry standards address

## Slide 69: M305 Supports
Transfer supports
Standing supports
Leg supports, stirrups
Head and back support

## Slide 70: Supports
Leg supports – must provide a method of supporting, positioning, and securing the patient’s legs
Head and back support – Where equipment is used in a reclined position, head and back support shall be provided. Where the incline of the back support can be modified while in use, it must be provided throughout the entire range of the incline.

## Slide 71: Transfer supports
Similar to ADA grab bars in section 609
Free of sharp or abrasive components
Eased edges
Shall not rotate within their fittings
Circular Cross Section
Outside diameter 1 ¼” min and 2” max
Non-circular Cross Section
Cross-section 2” max
Perimeter 4” min to 4.8” max

## Slide 72: Transfer supports
Gripping surfaces
Continuous along their length and not obstructed on top or sides
No more than 20% of the length of the bottom of gripping services can be obstructed
1 ½” minimum clearance

## Slide 73: Chapter 3
Supine, prone, or side-lying positions (M301)
Seated position (M302)
Seated in a wheelchair (M303)
Standing position (M304)
Supports (M305)
Communication (M306)
Operable parts (M307)
PUSH

## Slide 74: Communication (M306)
Information Communicated through the equipment
Information provided in at least two methods:
  audible
  visible
  tactile

## Slide 75: Chapter 3
PUSH
Supine, prone, or side-lying positions (M301)
Seated position (M302)
Seated in a wheelchair (M303)
Standing position (M304)
Supports (M305)
Communication (M306)
Operable parts (M307)

## Slide 76: Operable Parts (M307)
Only controls operated by the patient
Features addressed:
Tactilely Discernible
Operation
Operating Force

## Slide 77: DOJ Guidance is Available
www.ada.gov

## Slide 78: Final Standards
The final Standards for Accessible Medical Diagnostic Equipment can be found on www.access-board.gov under Healthcare, Medical Diagnostic Equipment

## Slide 79: Guidance on Prescription Drug Container Labels

## Slide 80: What Is a Prescription Drug Container Label
Legal Document
Universal Standard
U.S. Pharmacopeial Convention

## Slide 81: Section 904 Of the Food and Drug Administration Safety and Innovation Act
§904. ACCESSIBILITY OF INFORMATION ON PRESCRIPTION DRUG CONTAINER LABELS BY VISUALLY-IMPAIRED AND BLIND CONSUMERS.
Establishment of Working Group
Develop “best practices” on access to information on prescription drug container labels for individuals who are blind or visually impaired.

## Slide 82: Guidance on Prescription Drug Container Labels
(Advisory non-mandatory guidance)
Advisory guidance developed by a stakeholder working group provides best practices on providing access to prescription drug container labels for people who have vision impairments or who are elderly.

## Slide 83: National Council on Disability (NCD)
Section 904 of the Food and Drug Administration Safety and Innovation Act
The law directed the National Council on Disability (NCD) to conduct an informational and educational campaign in cooperation with the stakeholder working group to inform the public, including people with disabilities and pharmacists, of the best practices.
https://ncd.gov/publications/2016/best-practices-accessible-prescription-drug-labeling
The law also called upon the Comptroller General to conduct a review to assess the extent to which pharmacies are implementing the best practices and to determine whether barriers to prescription drug labels remain.
https://www.gao.gov/products/GAO-17-115

## Slide 84: Delivery Methods
Hard Copy Braille and Large Print
Digital Voice or Text to Speech Recorder
Radio Frequency Identification Device
Smart Devices and Computers

## Slide 85: Best Practices Recommendations
Best Practices for All Formats
Communication
Provide options
Quality control
Maintain patient privacy
No Surcharge (extra fee)
www.access-board.gov/guidelines-and-standards/health-care/about-prescription-drug-container-labels/working-group-recommendations

## Slide 86: Questions?

## Slide 87: Contact info
TECHNICAL ASSISTANCE
(800) 872-2253
TTY: (800) 993-2822
mde@access-board.gov
or
ta@access-board.gov
www.access-board.gov
Bobby Stinnette
(202) 272-0021
stinnette@access-board.gov

## Slide 88: Additional Resources
https://www.access-board.gov/guidelines-and-standards/health-care/about-this-rulemaking/final-standards/text-of-the-proposed-standards-8
https://adata.org/factsheet/accessible-medical-examination-tables-and-chairs
https://adata.org/factsheet/accessible-medical-diagnostic-equipment
https://www.ada.gov/medcare_mobility_ta/medcare_ta.htm
https://ncd.gov/publications/2016/best-practices-accessible-prescription-drug-labeling

## Slide 89: Summary: Supine, prone, or side-lying position (examination tables)
When adjustable, head and back support provided throughout the entire range of the incline (M301.3.3)
Transfer support resists vertical and horizontal forces of 250 lbs. at all points and does not rotate within its fittings (M305.2.2 and M305.2.3) Rail serves as transfer support within reach of transfer surface (M301.3.1 and M305.2.1)
When provided, stirrups provide a method of supporting, positioning, and securing the patient’s leg (M301.3.2)
Support rail removable / repositioned to permit unobstructed transfer (M301.2.3 EXCEPTION)
Transfer surface 30 inches wide minimum and 15 inches deep minimum (M301.2.2)
One short side (depth) and one long side (width) of the transfer surface permit unobstructed transfer from a mobility device (M301.2.3)
6 inches high minimum clearance above finished floor where equipment overhangs clearance (M301.4.1)
Base permits clearance around base for a patient portable floor lift, see Figure M2 (M301.4 and M301.4.2
Transfer surface 17 inches minimum and 19 inches maximum above floor level (M301.2.1), when not needed to facilitate transfer, the transfer surface may be positioned above or below the height range (Advisory M301.2.1)

## Slide 90: Summary: Seated position (examination chairs)
When adjustable, head and back support provided throughout the entire range of the incline (M302.3.3)
Transfer surface 21 inches wide minimum and 15 inches deep minimum (M302.2.2)
Armrest folds up to permit unobstructed transfer (M302.2.3 EXCEPTION)
Transfer surface 17 inches minimum and 19 inches maximum. Above floor (M302.2.1), when not needed to facilitate transfer, the transfer surface may be positioned above or below the height range (Advisory M302.2.1)
Required armrest serves as transfer support within reach of transfer surface (M302.3.1, M302.3.2, and M305.2.1)
Transfer support resists vertical and horizontal forces of 250 lbs. at all points and does not rotate within its fittings (M305.2.2 and M305.2.3)
One short side (depth) and one long side (width) of the transfer surface permit unobstructed transfer from a mobility device (M302.2.3)
Base permits clearance around base for a patient portable floor lift (M302.4 and M302.4.2)